## Slide 1: Elective: Developing Software for Automated MR QC Tests
Neil Heraghty

## Slide 2: Introduction
Project aim is to create an application that performs standard MRI QA measurements.
Desire is to perform tests with the minimum of user input.
Long-term aspiration for application to be web-based and free to use / contribute to.

## Slide 3: QC Tests

## Slide 4: IPEM Report 80 Tests
Uniformity
  Eurospin TO1A
  MagNET MAGFF-OIL
    Find images with a single edge?

## Slide 5: IPEM Report 80 Tests
SNR
  Eurospin TO1A
  MagNET MAGFF-LOADED
    Find images with a single edge?
      If/how to distinguish SNR/Uniformity images?

## Slide 6: IPEM Report 80 Tests
Ghosting
  Eurospin TO8
  Small water-based phantom
    Find images with a single edge?
      Object <25% of image area?

## Slide 7: IPEM Report 80 Tests
Spatial Resolution
  Eurospin TO4
  MagNET MAGRES
    Getting trickier...
      Look for edges?
      Squares?
      Cylindrical inserts could be used for MTF...

## Slide 8: IPEM Report 80 Tests
Geometric Distortion/Linearity
  Eurospin TO2
  MagNET MAGGEOM
    Hard to automate; requires prior knowledge of phantom.
      Phantom atlas?

## Slide 9: IPEM Report 80 Tests
Slice Width
  Eurospin TO2
  MagNET MAGGEOM
    Needs prior knowledge of plate angle...

## Slide 10: IPEM Report 80 Tests
Slice Position
  Eurospin TO3
  MagNET MAG_SLICEPOS
    More prior knowledge needed...
    Is the phantom set up correctly?

## Slide 11: IPEM Report 80 Tests
T1/T2 Relaxation Times
  Eurospin TO5
    Precise prior knowledge required – probably impossible to automate.

## Slide 12: Other tests
RF Interference
    Measurement of noise
      Not clear on the conditions of this test...

## Slide 13: Other tests
B0 Homogeneity
    Currently a qualitative test
      Not sure what this image shows...

## Slide 14: Other tests
What other tests might be desirable?

## Slide 15: Problems
Prior knowledge about test/phantom.
What to do when test requires multiple images (e.g. SNR, RF Interference, MAG_SLICEPOS)?

## Slide 16: Analysis of Daily QC

## Slide 17: Analysis of Daily QC
Routine QA with manufacturer phantoms

## Slide 19
SNR
Distortion
Ghosting
Uniformity

## Slide 21: The Plan – Part 1
Write Python scripts to test uniform phantoms
  Test using historical data from PACS
Develop system for automatically testing all daily/weekly flood phantom data
  Record results
  Alert users / physics if out of tolerance

## Slide 22: The Plan – Part 2
Develop web application to allow anyone to upload a DICOM image and perform the same tests
Add more tests / accepted test phantoms